## Slide 1: Development of a Test Facility for Studies Relevant to Replacing Dangerous Radiological Sources Long Vo
Aaron Hellinger, Maria Pinilla, William Dunn, Walter McNeil, Amir Bahadori
Department of Mechanical and Nuclear Engineering
Kansas State University
Manhattan, KS 66506 USA

## Slide 2: Preview
Motivation
Facility overview
Benchmarking tool
  Design
  D-T Generator
Test facility
  Location
  Design
  Operation
Representative test results
Summary

## Slide 3: Motivation
The KSU Materials Interrogation (KSUMI) Test Facility
  Replacement of Dangerous Radiological Sources (RDRS) – Current focus
  Detection and identification of Special Nuclear Materials (SNM)
  Radionuclide detection and identification in variable-background environments
RDRS is concerned with well logging tool:
  widely distributed to remote areas around the world
  contain multi-Ci neutron and gamma- ray sources such as 241Am-Be and 137Cs.
Chart: Total Crude Oil and Petroleum Products from USA
| 2012 | 2013 | 2014 | 2015 | 2016 | 2017
 | 6.7661 | 6.9229 | 6.9715 | 7.1297 | 7.2055 | 7.2551
Source: U.S. Geological Survey, February 2018

## Slide 4: NAS Study
The National Academy of Sciences (NAS) commissioned a study soon after 9/11/2001 to consider the dangers of large radiological sources
Radionuclide | Half-life (yr) | Particles-Energies (MeV) | Principal Applications | US Inventory (Ci)
Cs-137 | 30.17 | b-0.518 max; b-1.18 max; g-0.662 | Industrial gauging, Irradiators, Well logging | 2,800,000
Co-60 | 5.27 | g-1.173; g-1.333 | Sterilization, Irradiators, Teletherapy | 198,000,000
Am-241 | 432.2 | a-5.64; g-0.060 | Well logging, Gauging | 6,482
Pu-238 | 87.7 | a-5.59; g-0.043 | RTGs | 937
Sr-90 | 28.9 | b-0.546 max | RTGs | 1,730,000
Se-75 | 119.8 d | g-0.28 (ave), 0.8 max | Radiography | 261
Ir-192 | 74 d | b-1.46 max] g-0.38 ave, 1.378 max | Radiography | 146,922
Cf-252 | 2.645 | a-6.22; n-2 ave; g-various; fission products | Well logging Research | 7
Our focus on well logging is aimed at replacing 241Am-Be sources, which emit both neutrons and high-energy gamma rays, and 137Cs sources, which emit gamma rays and are usually in the form of water- soluble CsCl
We hope to show that a single deuterium-tritium (D-T) accelerator neutron source can replace both radioisotope sources
NAS findings

## Slide 5: Benchmarking Tool Design
The benchmarking tool:
  Skeleton
  Sources
  Shielding
  Sensors
For the well-logging application:
  241Am-Be radioisotope and Thermo Fisher D-T generator
  BF3 and 3He neutron sensors
  NaI(Tl) and CeBr3 gamma sensors
  Tool is used to benchmark Monte Carlo simulation results
BF3
Helium-3
NaI(Tl)
NaI(Tl)
Accelerator’s Target
Distances in inches, from target to centroid of active detection material
65.37’’
47.09’’
34.58’’
17.27’’
2’’ Lead
The Skeleton with only sensors

## Slide 6: D-T Generator: Upon Arrival
Arrival-Test for defects at Safety Office’s facility
Making
Connections
Established Communication
B322 arrived with 1631DT accelerator tube and was in possession of safety officers. Arrival-tests were done at safety office’s facility:
  Generator was filled to110 psi of 99.99% Sulfur Hexafluoride
  200 Volts D.C.
  Controlled via GUI by RS232 communication
Applications of the B322: oil & gas, mineral, environmental and imaging
Dimension: Cylindrical with 2 inches in diameter, 48 inches in length
Generator

## Slide 7: D-T Generator: GUI Controller
Interlock enabled GUI
80 kV, 60 uA = 1.6 x 108 nps determined by Copper activation
1 kHz 10% Duty Cycle
~45 seconds to the first neutron produced
~55 seconds to full power neutron production
GUI CONTROLLER

## Slide 8: Test Facility: Location
N
The KSUMI Test Facility is at King Hall Annex at KSU
North Side View
South Side View
Facility Floor Plan
Room 107
Room 101

## Slide 9: Test Facility: Accelerator Enabling and Safeguard Systems
Accelerator Enabling System prevents inadvertent neutron production and entries
Other Safeguards include:
  Surveillance
  Intercommunication
  Warning lights
  Procedures
Low dosage for bystanders and workers. Only workers who enters experiment room requires dosimeters.
Experimental Room
Storage
Cooler
Control Station
Source
D-T generator’s Target
Video & Audio Surveillance
Controlled Access
Hallway
N
Test Chamber

## Slide 10: Test Facility: Test Chamber Design
The test chamber is 6’-8” high by 6’-6” wide by 8’ deep
Volume inside the chamber is ~2,500 gal
When filled with water, the chamber presents an effectively infinite medium to the D-T neutrons
Tool goes inside the borehole tube which is surrounded by a medium such as air, water, sand, or limestone chips.
Test Chamber dimensions

## Slide 11: Test Facility: Test Chamber with Benchmark Tool
Green objects are borated polyethylene
Test Chamber is currently outfitted to benchmark the logging tool
Generator and tool inserting the chamber’s borehole tube
Shield Cap Interlock
Ready for measurements
Inside The Chamber
Borehole tube
½ ton crane
Open Borehole tube

## Slide 12: Test Facility: Control Room
Control Station
Surveillance display and Intercommunication
Measurements’ monitor and Control
Interlocks enabler
D-T generator’s monitor and GUI control

## Slide 13: Test Facility: Material Loading
Safety Evaluation
Conveyor Belt system for material movement
Test Chamber material leveling

## Slide 14: Test Facility: Data Acquisition Scheme
D-T Gen
G-far
G-near
N-near
N-far
PC
Ch1
Ch2
Ch3
Ch2
Ch1
Ch4
Ch0
Fiber Optic
CAEN5730 Digitizer
Physics
Data Acquisition Scheme
Output text options:
List Mode in ASCII format
List Mode in binary format
Function generator

## Slide 15: Neutron–Nucleus interaction
Neutron-Nucleus Interaction
Scattering
Absorption
Transfer
Elastic
(n,n)
Inelastic
(n,n’)
Capture
(n,γ)
fission
(n,f)
(n,2n)
(n,α)
(n,p)

## Slide 16: Representative Test Results: Water Sample
Events list
Triggers list
Expected Physics
Chart: Representative Result shows prompt and delay discrimination
Observed reactions
Observed prompt and delay gammas
Expected Physics: Die-away
Chart: 

## Slide 17: Potential Applications
The KSUMI test facility can be used to benchmark simulation models for oil well logging tools
We can conduct active interrogation with sources such as the D-T accelerator source or radioisotope sources such as 241Am-Be and 137Cs
We can also do passive interrogation by removing all sources from the tool
We can substitute other detectors, e.g. CeBr3 scintillation sensor and hand-held HPGe solid state sensor for remote source detection and possible identification

## Slide 18: ISOVOLT
Industrial X-ray Generator, 450 kV max
System constituents:
  X-ray Tube and Control Module
  Cooling Unit
  Digital Radiography flat panel imaging system
Currently configured for multiphase flow imaging
Operational constraint of 150 kV to limit human exposure

## Slide 19: Summary
The KSUMI capabilities: RDRS, detection of SNM, and source and material characterization studies
Full-scale test chamber and tool with variable sources, sensors, and shielding
So far, we have used both 241Am-Be and D-T sources, BF3 and 3He neutron detectors, and NaI(Tl) and CeBr3 photon detectors
Preliminary MCNP simulation models have been shown to agree with experimental results, which allows us to use the test facility as a benchmark and the simulation models to study materials and detectors that are difficult or expensive to use experimentally
This work is sponsored by NNSA through the Consortium for Nonproliferation Enabling Capabilities (CNEC)

## Slide 20: References
Bob Burkhart, “Neutron Generator and Well Logging”, Sandia National Laboratories, SAND2006-3291C, www.osti.gov/servlets/purl/1264595
Committee on Radiation Source and Replacement, 2008, Radiation Source Use and Replacement, National Research Council, ISBN 0-309-110115-7, National Academies Press, Washington, DC (NAS-2008)
Gardner, et al., 2013, ANS Panel on “Recent Developments in Radiation Source Use and Replacement after the NAS Report of 2008”, Transactions of the American Nuclear Society, Vol. 109, 47.
Liberman, A.D et. al., “A method to determine the absolute neutron output of small D-T generator”, Nuclear Instruments and Methods in Physics Research Section B, Volume 79, Issue 1-4, p. 574-578, June 1993
Preston R.M. et al, “Neutron generator burst timing measured using a pulse shape discrimination plastic scintillator with silicon photomultiplier readout”, Journal of Instrumentation, December 8, 2013
U.S. Total Crude Oil and Petroleum Products, U.S. Geological Survey, www.usgs.gov, accessed February 2018

## Slide 21: Thank you for your attention. Questions?
LONG VO
Email: Longv@ksu.edu